## bp
rule_no | rule_id | bp_section | question | on_yes | override_quid | override_rule | on_no | finding | recommendation
1 | p_1 | bp_philips | Has channel 144 been eliminated on APs supporting Philips devices? | PASS |  |  | FAIL | Channels 144 is available for use on APs suppporting Philips devices. Philips Healthcare devices do not support channel 144 and will be unable to associate with APs assigned this channel, causing the device to disconnect from the network | Remove channel 144 from use on APs supporting Philips Healthcare devices
2 | p_2 | bp_philips | Does RSSI meet Philips requirement of -67dBm or stronger? | PASS |  |  | FAIL | RSSI requirements for Philips Healthcare 802.11 networks do not meet the -67 dBm threshold in all areas | Reconfigure the network per recommedations within this report to acheive -67 dBm RSSI
3 | p_3 | bp_philips | Does SNR meet Philips requirement of 25dB or stronger? | PASS |  |  | FAIL | SNR requirements for Philips Healthcare 802.11 networks do not meet the 25 dB threshold in all areas | Reconfigure the network per recommedations within this report to acheive 25 dB SNR
4 | p_4 | bp_philips | Is primary channel separation a minimum of 20dB per Philips 802.11 requirements? | PASS |  |  | FAIL | Co-Channel separation does not meet the required level of 20 dB | Reconfigure channel assignment or AP power levels as recommended within this report
5 | p_5 | bp_philips | Are two or less APs on the primary channel at or above -82dBm per Philips 802.11 requirements? | PASS |  |  | FAIL | More than two APs on the primary channel at -82 dBm or stronger were detected over a majority of the scope area | Reconfigure channel assignment or AP power levels as recommended within this report
6 | p_6 | bp_philips | Is WPA2 or WPA2/3 Transistion Mode configured on APs supporting Philips devices? | PASS |  |  | FAIL | WPA3 only security is not compatible with Philips Healthcare devices. These devices will be unable to associate with the assigned SSID | Reconfigure SSID for WPA2 only or WPA2/3 Transition Mode
7 | p_7 | bp_philips | Is the SSID supporting Philips devices, broadcasting ? (Mandatory if DFS channels are in use) | PASS |  |  | FAIL | A Hidden SSID is not recommended and this SSID must be broadcast if DFS channels are available for use on APs supporting Philips Healthcare devices | Unhide (Broadcast) the SSID supporting Philips Healthcare devices
8 | bp_1_1 | bp1 | Are business critical applications limited to 5 GHz band with unique SSIDs? | PASS |  |  | FAIL | Business-critical applications are not limited to the 5 GHz band. This can cause slow performance and negatively impact roaming if 5 GHz capable clients connect to the 2.4 GHz band | SSIDs used by Life/Business critical applications should be configured on the 5 GHz band, only
9 | bp_1_2 | bp1 | Are Best Effort and IoT applications limited to 2.4 GHz band with unique SSIDs? | PASS |  |  | FAIL | Non-critical applications are not limited to the 2.4 GHz band. This can cause congestion on the 5 GHz band and degrade the performance of critical applications | SSIDs used by non-critical applications should be configured on the 2.4 GHz band, only
10 | bp_1_3 | bp1 | Are any SSIDs shared across both 2.4 and 5 GHz bands? | FAIL but can PASS | bp_1_3_justified | Is this justified? Justified if large number of devices originally setup on dual band SSID - If yes justified then this becomes a PASS. If No justified - FAIL | PASS | Identical SSIDs are broadcast on both the 2.4 and 5 GHz bands. This can cause 5 GHz capable devices to connect to the 2.4 GHz band, which may negatively impact performance and roaming | Create unique SSIDs for each band with a name that encourages 5 GHz capable clients to associate to the 5 GHz band (e.g., "SSID" for the 2.4 GHz band and "SSID-Best/Fast" for the 5 GHz band)
11 | bp_2_1 | bp2 | Is the 2.4 GHz network configured on non-overlapping channels 1, 6, and 11 only? | PASS |  |  | FAIL | Channels other than 1, 6, or 11 are in use on the 2.4 GHz band. These channels overlap the preferred channels and cause Adjacent Channel Interference (ACI). This can increase retries and CRC error rates | Reconfigure controller to allow only channels 1, 6, or 11 for use by access points
12 | bp_3_1 | bp3 | Are 5 GHz UNII-1 channels available for use and device compatibility verified? | PASS |  |  | FAIL | UNII-1 channels (36-48) are not available for use in the 5 GHz band. This limits the number of channels available for use by devices, which can increase Co-Channel Interference (CCI) | Enable UNII-1 channels for use by devices. Verify client devices are compatible with all channels selected in this band
13 | bp_3_2 | bp3 | Are 5 GHz UNII-2a channels available for use and device compatibility verified? | PASS |  |  | FAIL | UNII-2a channels (52-64) are not available for use in the 5 GHz band. This limits the number of channels available for use by devices, which can increase Co-Channel Interference (CCI) | Enable UNII-2a channels for use by devices. Verify client devices are compatible with all channels selected in this band. Monitor the system for DFS Channel Change events to determine if any of these channels are continually affected by radar. Remove specific channels within this band, if nescesary
14 | bp_3_3 | bp3 | Are 5 GHz UNII-2c channels available for use and device compatibility verified? | PASS |  |  | FAIL | UNII-2c channels (100-144) are not available for use in the 5 GHz band. This limits the number of channels available for use by devices, which can increase Co-Channel Interference (CCI) | Enable UNII-2c channels for use by devices. Verify client devices are compatible with all channels selected in this band. Monitor the system for DFS Channel Change events to determine if any of these channels are continually affected by radar. Remove specific channels within this band, if nescesary
15 | bp_3_4 | bp3 | Are 5 GHz UNII-3 channels available for use and device compatibility verified? | PASS |  |  | FAIL | UNII-3 channels (149-161) are not available for use in the 5 GHz band. This decreases the number of channels available for use by devices, which can increase Co-Channel Interference (CCI) | Enable UNII-3 channels for use by devices. Verify client devices are compatible with all channels selected in this band
16 | bp_3_5 | bp3 | Is Channel 165 (UNII-3) available for use and device compatibility verified? (20 MHz ) | PASS |  |  | FAIL | UNII- 3 channel 165 is not available for use in the 5 GHz band. This limits the number of channels available for use by devices, which can increase Co-Channel Interference (CCI) | Enable UNII-3 channel 165 for use by devices. Verify client devices are compatible with channel 165
19 | bp_4_1 | bp4 | Are 2.4 GHz Lowest Mandatory Rates limited to 12 and/or 24 Mbps? (1, 2, 5.5, 6, 9, 11 disabled) | PASS |  |  | FAIL | 2.4 GHz Lowest Mandatory data rates are not limited to 12 and/or 24 Mbps. These data rates are required to be supported by all wireless devices. | Reconfigure Mandatory data rates to include 12 and/or 24 Mbps, only
20 | bp_4_2 | bp4 | Are 5 GHz Lowest Mandatory Rates limited to 12 and/or 24 Mbps? (6/9 disabled) | PASS |  |  | FAIL | 5 GHz Mandatory data rates are not limited to 6, 12, and/or 24 Mbps. These data rates are required to be supported by all wireless devices | Reconfigure Mandatory data rates to include 12 and/or 24 Mbps, only
21 | bp_4_3 | bp4 | Is 6 Mbps eliminated from 5 GHz Mandatory Data Rates? | PASS | bp_4_3_justified | Is this justified? Justified if 12 Mbps is lowest Mandatory Rate. If justified then PASS | PASS | 6 Mbps data rate is set as a Mandatory data rate for the 5 GHz band. This can negatively impact performance | Disable 6 Mbps as a Mandatory Basic Rate
22 | bp_4_4 | bp4 | Is 12 Mbps eliminated from 5 GHz Mandatory Data Rates? | FAIL |  |  | PASS | 12 Mbps data rate is not set as a Mandatory data rate for the 5 GHz band but is recommended for this network. This can cause poor performance and disconnects | Reconfigure to allow 12 Mbps as a Mandatory Basic Rate
23 | bp_5_1 | bp5 | Are 20 MHz channels in use on the 2.4 GHz band? | PASS |  |  | FAIL | 40 MHz bonded channels are in use on the 2.4 GHz band. This limits the number of available channels and increases channel reuse | Reconfigure 2.4 GHz channels for 20 MHz wide channels, only
24 | bp_5_2 | bp5 | Are 20 MHz channels in use on the 5 GHz band? | PASS |  |  | FAIL | Bonded channels are in use on the 5 GHz band. The use of channel bonding reduces the number of available channels and increases channel reuse. Required data rates can be achieved utilizing lower channel bonding | Reconfigure 5 GHz channel width to 20 MHz wide channels, only
25 | bp_5_3 | bp5 | Are 40 MHz channels in use on the 5 GHz band? | FAIL but can PASS | bp_5_3_justified | Is this justified? Justified if required for higher throughput. If justified then PASS | PASS | 40 MHz bonded channels are in use on the 5 GHz band. The use of channel bonding reduces the number of available channels and increases channel reuse. Required data rates can be achieved utilizing lower channel bonding | Reconfigure 5 GHz channel width to 20 MHz wide channels, only
26 | bp_5_4 | bp5 | Are 80 MHz channels in use on the 5 GHz band? | FAIL |  |  | PASS | 80 MHz bonded channels are in use on the 5 GHz band. The use of channel bonding reduces the number of available channels and increases channel reuse. Required data rates can be achieved utilizing lower channel bonding | Reconfigure 5 GHz channel width to 20 MHz wide channels, only
27 | bp_5_5 | bp5 | Is DBS (Dynamic Bandwidth Selection) in use on the 5 GHz band? | FAIL |  |  | PASS | Dynamic Bandwidth Selection is in use. This can cause OBSS and roaming performance issues | Reconfigure Channel width to 20 MHz wide channels, only
29 | bp_6_1 | bp6 | Does 2.4 GHz RSSI meet target threshold of -72 dBm? | PASS |  |  | FAIL | Target RSSI requirement for clients or applications are not met. This can negatively impact performance | Reconfigure the network to acheive -72 dBm RSSI in all areas
30 | bp_6_2 | bp6 | Does 2.4 GHz SNR meet threshold of 20 dB? | PASS |  |  | FAIL | Target SNR requirement for clients or applications are not met. This can negatively impact performance | Reconfigure the network to acheive 20 dB SNR in all areas
31 | bp_7_1 | bp7 | Is co-channel interference nominal in 2.4 GHz band? | PASS |  |  | FAIL | Co-Channel Interference was detected in the 2.4 GHz band | Decrease power and/or disable 2.4 GHz radios as needed, to minimize CCI
32 | bp_8_1 | bp8 | Are 3 or less unique SSIDs in use on 2.4 GHz radios? | PASS |  |  | FAIL | More than 3 SSIDs are in use on the 2.4 GHz band. This increases overhead traffic and can cause poor performance | Reduce the number of SSIDs in use on 2.4 GHz radios. This can be accomplished by either eliminating SSIDs from use or assigning SSIDs only to radios in areas where they are required by devices
33 | bp_8_2 | bp8 | Are 4 or less unique SSIDs in use on 5 GHz radios? | PASS |  |  | FAIL | More than 4 SSIDs are in use on the 5 GHz band. This increases overhead traffic and can cause poor performance | Reduce the number of SSIDs in use on 5 GHz radios. This can be accomplished by either eliminating SSIDs from use or assigning SSIDs only to radios in areas where they are required by devices
34 | bp_8_3 | bp8 | Are all SSIDs in the WLAN being broadcast? (not hidden) | PASS | bp_8_3_justified | Is this justified? Justified if hidden SSID for use by stationary devices and does not require roaming. If justified then PASS | FAIL | Hidden SSIDs are in use on the network. This requires clients to probe for the network, instead of being able to discover it passively, which increases overhead traffic. SSIDs should only be hidden if in use by devices that are in static locations and do not require roaming | Reconfigure SSIDs to be broadcast, whenever possible
35 | bp_9_1 | bp9 | Is AES/CCMP encryption in use on all SSIDs? | PASS | bp_9_1_justified | Is this justified? Justified only if open guest networks are present. If justified then PASS. | FAIL | AES/CCMP encryption is not in use on all SSIDs. This can create security issues | Reconfigure security on all SSIDs to utilize AES/CCMP security
36 | bp_9_2 | bp9 | Is TKIP encryption in use on any SSIDs? (only justified if 54 Mbps max is acceptable) | FAIL but can PASS | bp_9_2_justified | Is this justified? Justified only if required by business critical devices. If justified then PASS | PASS | TKIP encryption is in use on SSIDs. This limits data rates to 54 Mbps and can cause poor network performance | Reconfigure SSID to eliminate TKIP
37 | bp_9_3 | bp9 | Is WEP encryption in use on any SSIDs? | FAIL |  |  | PASS | SSIDs exist with WEP. This can cause security issues | Reconfigure SSIDs to utilize WPA/WPA2 security
38 | bp_9_4 | bp9 | Is Pre-Shared Key (PSK) or 802.1x authentication in use on all SSIDs? | PASS |  |  | FAIL | Pre-Shared Key (PSK) or 802.1x authentication is not in use on all SSIDs. This can cause security issues | Reconfigure SSIDs to use PSK or 802.1x for authentication
39 | bp_9_5 | bp9 | Is an open SSID only used for Guest Wi-Fi access? | PASS |  |  | FAIL | Unsecured SSID is present. This can cause security issues | Reconfigure SSIDs to use PSK or 802.1x for authentication
40 | bp_9_6 | bp9 | Is 802.11r (Fast Transition) disabled OR configured in Adaptive Mode, only? | PASS |  |  | FAIL | 802.11r is enabled or in Non-Adaptive mode. This can cause issues with some devices that do not support it | Disable 802.11r or use 802.11r Adaptive Mode
41 | bp_10_1 | bp10 | Is the wireless controller on stable software? | PASS |  |  | FAIL | Wireless LAN Controller is not on a stable firmware version. This can result in poor performance and configuration of devices | Upgrade Wireless LAN Controller to the latest stable firmware that is compatible with the models of APs in use
42 | bp_10_2 | bp10 | Are APs properly mounted? | PASS |  |  | FAIL | APs are not properly mounted. Improper mounting can cause poor signal strength in parts of the intended coverage area and connectivity issues when clients move into those areas | Remount APs as close to center of intended coverage area, but at least 4 feet from obstructions or interferers
43 | bp_10_3 | bp10 | Are APs on the same software version as controller? | PASS |  |  | FAIL | APs are not on the same version as WLC. This can result in poor performance and configuration. | Upgrade AP images to same version as WLC, unless advised otherwise by controller manufacturer
44 | bp_10_4 | bp10 | Are AP power levels configured at or below client device capability? | PASS |  |  | FAIL | AP power levels are higher than client device capabilities. This can cause client devices to be unable to communicate with APs | Reconfigure AP power ranges to within client capabilities using RRM TPC or manual power assignments

## bp_headings
bp_id | order | heading | subheading
p1 | 0 | Philips Requirements
bp1 | 1 | Best Practice #1 - Band Usage | Make 5 GHz the primary band and use 2.4 GHz for best effort applications
bp2 | 2 | Best Practice #2 - 2.4 GHz Channels | Use only channels 1, 6, and 11 in the 2.4 GHz band
bp3 | 3 | Best Practice #3 - 5 GHz Channels | Use all 5 GHz channels that are supported by client devices
bp4 | 4 | Best Practice #4 - Basic Rates | Disable all lower rates unless required
bp5 | 5 | Best Practice #5 - Channel Bonding | Use bonded channels (40/80 MHz), when needed, in 5 GHz band only
bp6 | 6 | Best Practice #6 - Supported Rates | Base your target RSSI/SNR levels on client/application requirements
bp7 | 7 | Best Practice #7 - Interference | Reduce co-channel interference threshold to minimum possible levels
bp8 | 8 | Best Practice #8 - SSIDs | Limit the number of SSIDs per radio and justify any use of "hidden" SSIDs
bp9 | 9 | Best Practice #9 - Security | Avoid the use of TKIP, WEP, or None as encryption methods
bp10 | 10 | Best Practice #10 - Infrastructure | WLC and infrastructure support of network requirements

## bp (2)
Question # | bp_id | Philips Requirements | if ANSWER is YES | FAIL but can pass IF | if ANSWER IS NO | Finding | Recommendation
1 | p1 | Has channel 144 been eliminated on APs supporting Philips devices? | PASS |  | FAIL | Channels 144 is available for use on APs suppporting Philips devices. Philips Healthcare devices do not support channel 144 and will be unable to associate with APs assigned this channel, causing the device to disconnect from the network | Remove channel 144 from use on APs supporting Philips Healthcare devices
2 | p1 | Does RSSI meet Philips requirement of -67dBm or stronger? | PASS |  | FAIL | RSSI requirements for Philips Healthcare 802.11 networks do not meet the -67 dBm threshold in all areas | Reconfigure the network per recommedations within this report to acheive -67 dBm RSSI
3 | p1 | Does SNR meet Philips requirement of 25dB or stronger? | PASS |  | FAIL | SNR requirements for Philips Healthcare 802.11 networks do not meet the 25 dB threshold in all areas | Reconfigure the network per recommedations within this report to acheive 25 dB SNR
4 | p1 | Is primary channel separation a minimum of 20dB per Philips 802.11 requirements? | PASS |  | FAIL | Co-Channel separation does not meet the required level of 20 dB | Reconfigure channel assignment or AP power levels as recommended within this report
5 | p1 | Are two or less APs on the primary channel at or above -82dBm per Philips 802.11 requirements? | PASS |  | FAIL | More than two APs on the primary channel at -82 dBm or stronger were detected over a majority of the scope area | Reconfigure channel assignment or AP power levels as recommended within this report
6 | p1 | Is WPA2 or WPA2/3 Transistion Mode configured on APs supporting Philips devices? | PASS |  | FAIL | WPA3 only security is not compatible with Philips Healthcare devices. These devices will be unable to associate with the assigned SSID | Reconfigure SSID for WPA2 only or WPA2/3 Transition Mode
7 | p1 | Is the SSID supporting Philips devices, broadcasting ? (Mandatory if DFS channels are in use) | PASS |  | FAIL | A Hidden SSID is not recommended and this SSID must be broadcast if DFS channels are available for use on APs supporting Philips Healthcare devices | Unhide (Broadcast) the SSID supporting Philips Healthcare devices
Best Practice #1 - Band Usage
Make 5 GHz the primary band and use 2.4 GHz for best effort applications
8 | bp1 | Are business critical applications limited to 5 GHz band with unique SSIDs? | PASS |  | FAIL | Business-critical applications are not limited to the 5 GHz band. This can cause slow performance and negatively impact roaming if 5 GHz capable clients connect to the 2.4 GHz band | SSIDs used by Life/Business critical applications should be configured on the 5 GHz band, only
9 | bp1 | Are Best Effort and IoT applications limited to 2.4 GHz band with unique SSIDs? | PASS |  | FAIL | Non-critical applications are not limited to the 2.4 GHz band. This can cause congestion on the 5 GHz band and degrade the performance of critical applications | SSIDs used by non-critical applications should be configured on the 2.4 GHz band, only
10 | bp1 | Are any SSIDs shared across both 2.4 and 5 GHz bands? | FAIL but can PASS | Is this justified? Justified if large number of devices originally setup on dual band SSID - If yes justified then this becomes a PASS. If No justified - FAIL | PASS | Identical SSIDs are broadcast on both the 2.4 and 5 GHz bands. This can cause 5 GHz capable devices to connect to the 2.4 GHz band, which may negatively impact performance and roaming | Create unique SSIDs for each band with a name that encourages 5 GHz capable clients to associate to the 5 GHz band (e.g., "SSID" for the 2.4 GHz band and "SSID-Best/Fast" for the 5 GHz band)
Best Practice #2 - 2.4 GHz Channels
Use only channels 1, 6, and 11 in the 2.4 GHz band
11 | bp2 | Is the 2.4 GHz network configured on non-overlapping channels 1, 6, and 11 only? | PASS |  | FAIL | Channels other than 1, 6, or 11 are in use on the 2.4 GHz band. These channels overlap the preferred channels and cause Adjacent Channel Interference (ACI). This can increase retries and CRC error rates | Reconfigure controller to allow only channels 1, 6, or 11 for use by access points
Best Practice #3 - 5 GHz Channels
Use all 5 GHz channels that are supported by client devices
12 | bp3 | Are 5 GHz UNII-1 channels available for use and device compatibility verified? | PASS |  | FAIL | UNII-1 channels (36-48) are not available for use in the 5 GHz band. This limits the number of channels available for use by devices, which can increase Co-Channel Interference (CCI) | Enable UNII-1 channels for use by devices. Verify client devices are compatible with all channels selected in this band
13 | bp3 | Are 5 GHz UNII-2a channels available for use and device compatibility verified? | PASS |  | FAIL | UNII-2a channels (52-64) are not available for use in the 5 GHz band. This limits the number of channels available for use by devices, which can increase Co-Channel Interference (CCI) | Enable UNII-2a channels for use by devices. Verify client devices are compatible with all channels selected in this band. Monitor the system for DFS Channel Change events to determine if any of these channels are continually affected by radar. Remove specific channels within this band, if nescesary
14 | bp3 | Are 5 GHz UNII-2c channels available for use and device compatibility verified? | PASS |  | FAIL | UNII-2c channels (100-144) are not available for use in the 5 GHz band. This limits the number of channels available for use by devices, which can increase Co-Channel Interference (CCI) | Enable UNII-2c channels for use by devices. Verify client devices are compatible with all channels selected in this band. Monitor the system for DFS Channel Change events to determine if any of these channels are continually affected by radar. Remove specific channels within this band, if nescesary
15 | bp3 | Are 5 GHz UNII-3 channels available for use and device compatibility verified? | PASS |  | FAIL | UNII-3 channels (149-161) are not available for use in the 5 GHz band. This decreases the number of channels available for use by devices, which can increase Co-Channel Interference (CCI) | Enable UNII-3 channels for use by devices. Verify client devices are compatible with all channels selected in this band
16 | bp3 | Is Channel 165 (UNII-3) available for use and device compatibility verified? (20 MHz ) | PASS |  | FAIL | UNII- 3 channel 165 is not available for use in the 5 GHz band. This limits the number of channels available for use by devices, which can increase Co-Channel Interference (CCI) | Enable UNII-3 channel 165 for use by devices. Verify client devices are compatible with channel 165
Best Practice #4 - Basic Rates
Disable all lower rates unless required
19 | bp4 | Are 2.4 GHz Lowest Mandatory Rates limited to 12 and/or 24 Mbps? (1, 2, 5.5, 6, 9, 11 disabled) | PASS |  | FAIL | 2.4 GHz Lowest Mandatory data rates are not limited to 12 and/or 24 Mbps. These data rates are required to be supported by all wireless devices. | Reconfigure Mandatory data rates to include 12 and/or 24 Mbps, only
20 | bp4 | Are 5 GHz Lowest Mandatory Rates limited to 12 and/or 24 Mbps? (6/9 disabled) | PASS |  | FAIL | 5 GHz Mandatory data rates are not limited to 6, 12, and/or 24 Mbps. These data rates are required to be supported by all wireless devices | Reconfigure Mandatory data rates to include 12 and/or 24 Mbps, only
21 | bp4 | Is 6 Mbps eliminated from 5 GHz Mandatory Data Rates? | PASS | Is this justified? Justified if 12 Mbps is lowest Mandatory Rate. If justified then PASS | FAIL | 6 Mbps data rate is set as a Mandatory data rate for the 5 GHz band. This can negatively impact performance | Disable 6 Mbps as a Mandatory Basic Rate
22 | bp4 | Is 12 Mbps eliminated from 5 GHz Mandatory Data Rates? | FAIL |  | PASS | 12 Mbps data rate is not set as a Mandatory data rate for the 5 GHz band but is recommended for this network. This can cause poor performance and disconnects | Reconfigure to allow 12 Mbps as a Mandatory Basic Rate
Best Practice #5 - Channel Bonding
Use bonded channels (40/80 MHz), when needed, in 5 GHz band only
23 | bp5 | Are 20 MHz channels in use on the 2.4 GHz band? | PASS |  | FAIL | 40 MHz bonded channels are in use on the 2.4 GHz band. This limits the number of available channels and increases channel reuse | Reconfigure 2.4 GHz channels for 20 MHz wide channels, only
24 | bp5 | Are 20 MHz channels in use on the 5 GHz band? MISSING IN LIST | PASS | Is this justified? Justified if required for higher throughput. If justified then PASS | FAIL | Bonded channels are in use on the 5 GHz band. The use of channel bonding reduces the number of available channels and increases channel reuse. Required data rates can be achieved utilizing lower channel bonding | Reconfigure 5 GHz channel width to 20 MHz wide channels, only
25 | bp5 | Are 40 MHz channels in use on the 5 GHz band? | FAIL but can PASS | Is this justified? Justified if required for higher throughput. If justified then PASS | PASS | 40 MHz bonded channels are in use on the 5 GHz band. The use of channel bonding reduces the number of available channels and increases channel reuse. Required data rates can be achieved utilizing lower channel bonding | Reconfigure 5 GHz channel width to 20 MHz wide channels, only
26 | bp5 | Are 80 MHz channels in use on the 5 GHz band? | FAIL |  | PASS | 80 MHz bonded channels are in use on the 5 GHz band. The use of channel bonding reduces the number of available channels and increases channel reuse. Required data rates can be achieved utilizing lower channel bonding | Reconfigure 5 GHz channel width to 20 MHz wide channels, only
27 | bp5 | Is DBS (Dynamic Bandwidth Selection) in use on the 5 GHz band? | FAIL |  | PASS | Dynamic Bandwidth Selection is in use. This can cause OBSS and roaming performance issues | Reconfigure Channel width to 20 MHz wide channels, only
Best Practice #6 - Supported Rates
Base your target RSSI/SNR levels on client/application requirements
29 | bp6 | Does 2.4 GHz RSSI meet target threshold of -72 dBm? | PASS |  | FAIL | Target RSSI requirement for clients or applications are not met. This can negatively impact performance | Reconfigure the network to acheive -72 dBm RSSI in all areas
30 | bp6 | Does 2.4 GHz SNR meet threshold of 20 dB? | PASS |  | FAIL | Target SNR requirement for clients or applications are not met. This can negatively impact performance | Reconfigure the network to acheive 20 dB SNR in all areas
Best Practice #7 - Interference
Reduce co-channel interference threshold to minimum possible levels
31 | bp7 | Is co-channel interference nominal in 2.4 GHz band? | PASS |  | FAIL | Co-Channel Interference was detected in the 2.4 GHz band | Decrease power and/or disable 2.4 GHz radios as needed, to minimize CCI
Best Practice #8 - SSIDs
Limit the number of SSIDs per radio and justify any use of "hidden" SSIDs
32 | bp8 | Are 3 or less unique SSIDs in use on 2.4 GHz radios? | PASS |  | FAIL | More than 3 SSIDs are in use on the 2.4 GHz band. This increases overhead traffic and can cause poor performance | Reduce the number of SSIDs in use on 2.4 GHz radios. This can be accomplished by either eliminating SSIDs from use or assigning SSIDs only to radios in areas where they are required by devices
33 | bp8 | Are 4 or less unique SSIDs in use on 5 GHz radios? | PASS |  | FAIL | More than 4 SSIDs are in use on the 5 GHz band. This increases overhead traffic and can cause poor performance | Reduce the number of SSIDs in use on 5 GHz radios. This can be accomplished by either eliminating SSIDs from use or assigning SSIDs only to radios in areas where they are required by devices
34 | bp8 | Are all SSIDs in the WLAN being broadcast? (not hidden) | PASS | Is this justified? Justified if hidden SSID for use by stationary devices and does not require roaming. If justified then PASS | FAIL | Hidden SSIDs are in use on the network. This requires clients to probe for the network, instead of being able to discover it passively, which increases overhead traffic. SSIDs should only be hidden if in use by devices that are in static locations and do not require roaming | Reconfigure SSIDs to be broadcast, whenever possible
Best Practice #9 - Security
Avoid the use of TKIP, WEP, or None as encryption methods
35 | bp9 | Is AES/CCMP encryption in use on all SSIDs? | PASS | Is this justified? Justified only if open guest networks are present. If justified then PASS. | FAIL | AES/CCMP encryption is not in use on all SSIDs. This can create security issues | Reconfigure security on all SSIDs to utilize AES/CCMP security
36 | bp9 | Is TKIP encryption in use on any SSIDs? (only justified if 54 Mbps max is acceptable) | FAIL but can PASS | Is this justified? Justified only if required by business critical devices. If justified then PASS | PASS | TKIP encryption is in use on SSIDs. This limits data rates to 54 Mbps and can cause poor network performance | Reconfigure SSID to eliminate TKIP
37 | bp9 | Is WEP encryption in use on any SSIDs? MISSING IN LIST | FAIL |  | PASS | SSIDs exist with WEP. This can cause security issues | Reconfigure SSIDs to utilize WPA/WPA2 security
38 | bp9 | Is Pre-Shared Key (PSK) or 802.1x authentication in use on all SSIDs? | PASS |  | FAIL | Pre-Shared Key (PSK) or 802.1x authentication is not in use on all SSIDs. This can cause security issues | Reconfigure SSIDs to use PSK or 802.1x for authentication
39 | bp9 | Is an open SSID only used for Guest Wi-Fi access? | PASS |  | FAIL | Unsecured SSID is present. This can cause security issues | Reconfigure SSIDs to use PSK or 802.1x for authentication
40 | bp9 | Is 802.11r (Fast Transition) disabled OR configured in Adaptive Mode, only? | PASS |  | FAIL | 802.11r is enabled or in Non-Adaptive mode. This can cause issues with some devices that do not support it | Disable 802.11r or use 802.11r Adaptive Mode
Best Practice #10 - Infrastructure
WLC and infrastructure support of network requirements
41 | bp10 | Is the wireless controller on stable software? | PASS |  | FAIL | Wireless LAN Controller is not on a stable firmware version. This can result in poor performance and configuration of devices | Upgrade Wireless LAN Controller to the latest stable firmware that is compatible with the models of APs in use
42 | bp10 | Are APs properly mounted? | PASS |  | FAIL | APs are not properly mounted. Improper mounting can cause poor signal strength in parts of the intended coverage area and connectivity issues when clients move into those areas | Remount APs as close to center of intended coverage area, but at least 4 feet from obstructions or interferers
43 | bp10 | Are APs on the same software version as controller? | PASS |  | FAIL | APs are not on the same version as WLC. This can result in poor performance and configuration. | Upgrade AP images to same version as WLC, unless advised otherwise by controller manufacturer
44 | bp10 | Are AP power levels configured at or below client device capability? | PASS |  | FAIL | AP power levels are higher than client device capabilities. This can cause client devices to be unable to communicate with APs | Reconfigure AP power ranges to within client capabilities using RRM TPC or manual power assignments